AeroCardia — Week 3 Final Findings

5 Datasets Analysed | Updated for Team Meeting

Analyst: Kenisha  |  Branch: kenisha_dev  |  AeroCardia DS Internship  |  Week 3

# 1. All Sessions Analysed — Summary

Five datasets were received and analysed. Results reveal a consistent pattern: three sessions belong to the same patient who has never produced a complete usable clinical session.

| Dataset | Duration | Rows | Patient | Status | Clinical Data |
| --- | --- | --- | --- | --- | --- |
| D1 — TCC | 35.9s | 360 | Patient A | Calibration | Partial PPG only — artefact study |
| D2 — Tracker | 134s | 6,007 | Patient B | No data | All sensors zero. O2 = 8% (impossible) |
| D3 — User session | 12.5 min | 26,312 | Patient C | Real data | HR, SpO2 active. HRV placeholder. CO2/LV inactive. |
| D4 — Tracker | 34.4s | 1,615 | Patient B | No data | PPG completely off. All clinical sensors zero. |
| D5 — Tracker | 23.6s | 1,096 | Patient B | Partial/failed | Brief HR lock-on (6.5s). O2=8% impossible. SpO2 locked at 99. |

# 2. Critical Finding — Repeat Patient Failure Pattern

Patient B (ID: 3c9674b4) appears in three datasets — D2, D4, and D5. All three sessions have failed to produce usable clinical data. This is the most important finding from the full dataset analysis.

| Session | Duration | HR | SpO2 | Key Issue |
| --- | --- | --- | --- | --- |
| D2 | 134s | Zero | Zero | All PPG sensors off. O2 = 8% (device not on body) |
| D4 | 34.4s | Zero | Zero | PPG completely inactive throughout entire session |
| D5 | 23.6s | 6.5s only | Locked 99 | Brief lock-on (107-112 bpm). O2 = 8% (impossible). SpO2 only reads 0 or 99 — not real. |

What this tells us:

- Patient B has attempted the device three times without a successful session

- The O2 reading of 8-9% in D2 and D5 is physically impossible for ambient air — this strongly suggests the device was not on the body when the session started

- In D5 the PPG LED (raw signal) IS firing, but the HR algorithm could not lock on — this points to incorrect device placement or movement artefact rather than a hardware failure

- This patient likely needs direct support on how to correctly wear and position the device before starting a session

- Recommendation: AeroCardia should implement a real-time skin-contact confirmation before allowing a session to begin

# 3. Sensor Activation — All 5 Datasets

| Sensor | D1 | D2 | D3 | D4 | D5 |
| --- | --- | --- | --- | --- | --- |
| Heart Rate | Artefact | Inactive | Active | Inactive | 6.5s only |
| SpO2 | Active | Inactive | Active* | Inactive | Locked 99 |
| HRV SDNN | Inactive | Inactive | Placeholder | Inactive | Inactive |
| O2 | Normal | Impossible | Normal | Normal | Impossible |
| CO2 | Inactive | Inactive | Inactive | Inactive | Inactive |
| Lung Volume | Placeholder | Inactive | Inactive | Inactive | Inactive |
| PPG Raw Signal | Partial | Inactive | Partial | Inactive | Active |
| Pressure sensors | Active | Active | Active | Active | Active |
| Temperature | Active | Partial | Active | Active | Active |
| Battery reporting | Inactive | Inactive | Inactive | Inactive | Active! |

* D3 SpO2 shows 372 readings below 95% in the real signal — AeroCardia cap rule technically triggered. Clarification needed on warm-up exclusion.

# 4. Dataset 5 — New Findings

## 4a. PPG LED Active But HR Algorithm Cannot Lock On

This is a new finding not seen in previous datasets. In D5 the raw PPG signals (ppg_ir and ppg_red) are continuously active throughout the session — the LED is firing and collecting light data. However the HR computation algorithm only managed to calculate a heart rate for the final 6.5 seconds (298 out of 1096 rows, 27% of session).

This means the problem is not a hardware failure — the sensor is physically working. The issue is that the algorithm requires a stable, consistent signal to compute HR and is not getting one. Most likely causes:

- Device moving on the body — the PPG signal is very sensitive to motion

- Incorrect placement — not pressed firmly enough against skin

- Session started before device was positioned on body

## 4b. SpO2 Locked at 99% — Not Real Data

In the 298 rows where a reading exists, SpO2 is exactly 99% for every single row. Real SpO2 varies slightly beat-to-beat. A locked value of 99 with no variation is a sensor initialisation lock-on, not a real measurement. Combined with the short 23-second session this value should not be used clinically.

## 4c. O2 at 8.5% — Device Not on Body

O2 reads 8.5% throughout the entire session. This is physically impossible in normal air (20.9% O2). The same pattern appeared in Dataset 2. The most likely explanation is that the O2 sensor was enclosed or covered in a way that depleted oxygen near the sensor — consistent with the device sitting face-down on a surface or inside a case rather than being worn on a body.

## 4d. Battery State = 1 — First Non-Zero Reading

Dataset 5 is the first session across all five datasets where battery_state reports a non-zero value (value = 1). This suggests battery reporting may now be partially active in a newer firmware version. Worth flagging to AeroCardia to confirm what value 1 represents.

# 5. Pipeline Bottlenecks — Consistent Across All Sessions

| Bottleneck | Severity | Evidence |
| --- | --- | --- |
| No complete protocol session | Critical | Longest session = 12.5 min (D3). Minimum protocol = 15 min. Zero domain scores computable. |
| CO2 sensor inactive | Critical | All zeros in every single dataset. Blocks RER and respiratory efficiency analysis. |
| Lung Volume not measured | Critical | Inactive or placeholder in all 5 datasets. Blocks FEV1, FVC, FEV1/FVC — 70% of Pulmonary Score. |
| HRV firmware issue | Critical | D3: only 2 unique values (427.9, 439.3ms). Normal = 20-100ms. Not real variability data. |
| PPG reliability | High | PPG inactive or artefact-only in 4 out of 5 sessions. Root cause: device placement. |
| Patient B wear compliance | High | 3 failed sessions from same patient. O2=8% confirms device not on body at session start. |
| No skin contact check | High | App allows sessions to start without confirming device is on skin. Causes silent data loss. |
| No minimum session enforcement | Medium | Sessions as short as 23s are recorded as source=full. App needs a minimum time gate. |
| SpO2 warm-up artefact | Medium | D3: 189 rows at 89% SpO2 during warm-up. Cap rule ambiguity — spec does not address this. |

# 6. Device Improvement Recommendations

Based on analysis of all five sessions, the following improvements would most significantly impact consumer satisfaction and data quality. Prioritised by impact:

## Critical — Must fix before consumer launch

- Real-time skin contact confirmation: App must verify PPG sensor detects skin before allowing session to start. Three out of five sessions failed because device was not properly worn.

- CO2 and Lung Volume activation: Both sensors are listed as core features in the spec but have never produced data across any session. These enable spirometry scoring — the largest component of the Pulmonary Score.

- HRV firmware fix: HRV SDNN is returning only 2 values across an entire 12.5-minute session. Real HRV varies continuously. This needs to be recomputed correctly before the metric is usable.

- Minimum session enforcement: The app records sessions as short as 23 seconds as source=full. A minimum 15-minute gate should prevent incomplete sessions from being saved as complete.

## Important — Fix before launch

- Live sensor status feedback: Users need real-time confirmation that HR and SpO2 are being detected during a session. Silent zero readings cause confusion and data loss.

- Session failure notification: When a session ends with no clinical data, the user should be told clearly — not shown blank results.

- SpO2 warm-up exclusion rule: The spec does not address whether warm-up period SpO2 readings count toward the 95% cap rule. This needs a clinical decision from AeroCardia.

- Battery status: Battery reporting is zero in 4 of 5 datasets. Now partially active in D5. Full implementation needed so users are warned before a session dies mid-test.

## Nice to have — Post-launch improvements

- Device placement guide in app: Step-by-step visual guidance on correct placement before each session starts.

- User profile linkage: Sessions are currently anonymous — no age, sex, height, weight attached. Without this, normative scoring against the spec cannot run.

- Session quality score: After each session show the user a simple quality indicator — how many sensors were active, how much data was captured — so they know if the session was valid.

# 7. Questions for AeroCardia — Today's Meeting

- Patient B has three failed sessions with O2 reading 8-9%. Is this a known firmware issue or a device wear problem? What support can be offered to this user?

- In Dataset 5 the PPG LED is active but HR algorithm only locked on for 6.5 seconds. What signal quality threshold does the HR algorithm require to compute a reading?

- CO2 and Lung Volume have been inactive across all five datasets. What is the timeline for activating these sensors in the firmware?

- HRV SDNN shows only two values in the real user session. Is this a known firmware bug? When will it be corrected?

- Should warm-up period SpO2 readings be excluded from the Pulmonary Score cap rule?

- When can we expect a complete 15 or 20-minute protocol session with all sensors active?

- Battery state = 1 appears in Dataset 5 for the first time. What does value 1 represent — is this a new firmware version?

# 8. Files Committed to kenisha_dev

| File | Description |
| --- | --- |
| aerocardia_tcc_cleaned.csv | Cleaned Dataset 1 — TCC calibration (360 rows) |
| aerocardia_user_session_cleaned.csv | Cleaned Dataset 3 — Real user session (26,312 rows) |
| aerocardia_dataset4_cleaned.csv | Cleaned Dataset 4 — Failed session, Patient B (1,615 rows) |
| aerocardia_dataset5_cleaned.csv | Cleaned Dataset 5 — Partial session, Patient B (1,096 rows) |
| kenisha_dev_findings_v3.docx | This document — Week 3 final findings across all 5 datasets |

AeroCardia DS Internship  |  Week 3 Final Findings  |  Analyst: Kenisha  |  Branch: kenisha_dev  |  5 Datasets Analysed