AeroCardia — Initial Findings & Observations

Week 3 Deliverable — Updated with Real User Session

Analyst: Kenisha  |  Branch: kenisha_dev  |  Datasets analysed: 3  |  Prepared: May 2026

# 1. Datasets Analysed — Overview

Three datasets were received and analysed during Week 3. Only one contains a real user session with active clinical sensors.

| Dataset | Duration | Status | Rows | Usable? |
| --- | --- | --- | --- | --- |
| Dataset 1 — TCC calibration | 35.9 sec | Calibration | 360 | Partial — artefact study only |
| Dataset 2 — Tracker (no data) | 2.2 min | Failed | 6,007 | No — all sensors at zero |
| Dataset 3 — Real user session | 12.5 min | Real data | 26,312 | Yes — primary analysis dataset |

# 2. Sensor Activation Comparison

This table shows which sensors were active across all three datasets. The progression reveals the device firmware becoming more active over time.

| Sensor | Dataset 1 | Dataset 2 | Dataset 3 |
| --- | --- | --- | --- |
| Heart Rate | Artefact only | Inactive | Active |
| SpO2 | Active (96-98%) | Inactive | Active (flag noted) |
| HRV SDNN | Inactive | Inactive | Placeholder values |
| O2 | Declining trend | 8% (impossible) | Normal 21% |
| CO2 | Inactive | Inactive | Inactive |
| Lung Volume | Placeholder 4.5L | Inactive | Inactive |
| PPG IR/Red | Active | Inactive | Partial |
| Pressure sensors | Active | Active | Active |
| Temperature | Active | Partial | Active |

# 3. Data Cleaning Summary

Each dataset was cleaned separately. All rows were retained — nothing was deleted. Issues were handled using flag columns so the original data is preserved.

## Dataset 1 — TCC Calibration Session

- Dropped 5 columns: HRV SDNN (100% NaN), CO2 (all zeros), RER (constant 0.7), Lung Volume (constant 4.5L), Battery State (all zeros)

- Flagged 30 rows as PPG warm-up (HR and SpO2 NaN during sensor initialisation)

- Flagged 21 rows as HR lock-on artefact (HR=120 before real signal acquired)

- Created heart_rate_clean_bpm column: 309 valid readings, range 49-78 bpm

- All columns renamed to snake_case for repository compatibility

## Dataset 2 — Failed Session (no data)

- All clinical sensors at zero throughout — heart rate, SpO2, PPG all inactive

- O2 reading of 8-9% is physically impossible (ambient air is ~20.9%)

- Tracker status confirmed as "no data" — session failed, likely device not worn correctly

- Not used for clinical analysis — documented as a pipeline failure case

## Dataset 3 — Real User Session (primary dataset)

- Dropped 2 columns: Lung Volume (100% NaN), CO2 (all zeros)

- Flagged 530 rows as PPG warm-up (first ~15 seconds, HR and SpO2 at zero)

- Flagged 38 rows as HR lock-on artefact (HR=127 spike at 16-17 seconds)

- Flagged 189 rows as SpO2 warm-up drop (SpO2 drops to 89% at 23-28 seconds — sensor artefact)

- Flagged all 26,312 rows as HRV placeholder (only 2 unique values — not real variability data)

- Created clean columns: heart_rate_clean_bpm (25,744 valid), spo2_clean_pct, hrv_sdnn_clean_ms

- All columns renamed to snake_case

# 4. Clinical Signal Analysis — Dataset 3 (Real User Session)

## 4a. Heart Rate

| Metric | Value | Notes |
| --- | --- | --- |
| Valid clean readings | 25,744 rows | Out of 26,312 total |
| Range (clean) | 51-126 bpm | Physiologically normal |
| Mean | 79.5 bpm | Consistent with light-moderate activity |
| Std deviation | 6.7 bpm | Low variability — stable signal |
| Warm-up NaN rows | 530 rows | First 15 seconds — expected PPG behaviour |
| Artefact rows (HR=127) | 38 rows | 16-17 seconds — lock-on spike, flagged |

HR Trend — Clear exercise response pattern observed:

| Time segment | Mean HR | Interpretation |
| --- | --- | --- |
| 0.0-2.5 minutes | 73.4 bpm | Resting / early session baseline |
| 2.5-5.0 minutes | 80.7 bpm | Rising — physical test beginning |
| 5.0-7.5 minutes | 81.0 bpm | Plateau — sustained activity |
| 7.5-10.0 minutes | 81.0 bpm | Plateau continues |
| 10.0-12.5 minutes | 81.0 bpm | Stable — consistent exercise state |

This progressive rise from 73 to 81 bpm followed by a stable plateau is a classic physiological exercise response. The user was likely performing one of the AeroCardia physical tests during this session.

## 4b. SpO2 — Clinical Flag Triggered

| Metric | Value | Notes |
| --- | --- | --- |
| Valid clean readings | 25,593 rows | After warm-up exclusion |
| Normal range (clean) | 95-99% | Majority of session is normal |
| Warm-up drop (artefact) | 89% at 23-28s | 189 rows — sensor initialisation, flagged separately |
| Readings below 95% (real signal) | 372 rows | After warm-up excluded — AeroCardia cap rule applies |
| AeroCardia cap rule | TRIGGERED | Pulmonary Score would be capped at 40 per spec |

IMPORTANT QUESTION FOR AEROCARDIA: The spec states SpO2 < 95% caps the Pulmonary Score at 40. The 189-row drop at 23-28 seconds appears to be a sensor warm-up artefact. However there are also 372 readings below 95% in the real signal portion of the session. AeroCardia should clarify: (1) should warm-up SpO2 readings be excluded from the cap rule, and (2) what is the clinical significance of the 372 real readings below 95%?

## 4c. HRV SDNN — Firmware Issue Identified

| Metric | Value | Notes |
| --- | --- | --- |
| Unique values in entire session | 2 only | 427.9 ms and 439.3 ms |
| Normal healthy adult SDNN | 20-100 ms | Per Nunan 2010 (AeroCardia spec reference) |
| Assessment | NOT real data | Only 2 values, both >400ms — physiologically implausible |
| Usable HRV readings | 0 | All 26,312 rows flagged as placeholder |

HRV SDNN is listed as a key metric in the AeroCardia spec (10% weight in Pulmonary Score, 20/30-min protocol). The firmware appears to be computing it incorrectly or returning a default value. This needs to be raised with AeroCardia as a pipeline bottleneck.

## 4d. O2 — Normal in Dataset 3

O2 reads 20.8-21.2% throughout the session, consistent with ambient air (~20.9%). This is normal and expected. Contrast this with Dataset 2 where O2 read 8-9% — confirming Dataset 2 was a failed/non-worn session.

# 5. Pipeline Bottlenecks — Across All Sessions

| Bottleneck | Severity | Detail |
| --- | --- | --- |
| CO2 sensor inactive | High | Inactive in all 3 datasets. Needed for RER calculation. |
| Lung Volume not measured | High | Inactive in all 3 datasets. Blocks FEV1 and FVC scoring entirely. |
| HRV firmware error | High | Only 2 unique values (427.9, 439.3 ms) across full 12.5-min session. Not real data. |
| No complete protocol session | High | Longest session is 12.5 min. Shortest AeroCardia protocol is 15 min. No domain scores computable yet. |
| SpO2 cap rule ambiguity | Medium | Spec does not distinguish warm-up artefact from real low SpO2. Clarification needed. |
| Session failures (Dataset 2) | Medium | One session produced zero clinical data. Device wear compliance may be an issue. |
| Battery not reporting | Low | All zeros across all datasets. Operational concern only. |

# 6. Key Observations

- Dataset 3 is the first real user session received. It shows a clear physiological exercise response in HR — rising from 73 bpm at rest to a stable plateau of 81 bpm — which is clinically meaningful and consistent with a physical test being performed.

- The most critical finding is the HRV firmware issue. HRV carries 10% weight in the Pulmonary Score per the AeroCardia spec. The current firmware appears to return placeholder values rather than computing real heart rate variability. This blocks a key scoring component.

- CO2 and Lung Volume have been inactive across all three datasets. Without these, spirometry-based scoring (FEV1, FVC, FEV1/FVC — which together carry 70% of the Pulmonary Score) cannot be computed.

- Dataset 2 with all sensors at zero and O2 at 8% is a confirmed failed session — the device was either not worn or not initialised. This is a user compliance or device setup issue worth flagging to AeroCardia.

- The PPG warm-up and HR lock-on artefact patterns are consistent across Datasets 1 and 3. This is expected device behaviour and has been correctly identified and flagged in both cleaned files.

- No complete AeroCardia protocol session (minimum 15 minutes) has been received yet. Full domain scoring cannot begin until a complete session is available.

# 7. Questions for AeroCardia — Meeting Agenda

- Can you confirm whether Dataset 1 (36 seconds) is a calibration run rather than a real user session?

- The HRV SDNN column shows only two values (427.9 and 439.3 ms) in the 12.5-minute session. Normal SDNN is 20-100 ms. Is this a firmware computation issue?

- SpO2 drops to 89% during the warm-up period (23-28 seconds) in Dataset 3. Should warm-up SpO2 readings be excluded from the Pulmonary Score cap rule?

- CO2 and Lung Volume are inactive across all three datasets. When will these sensors be active in sessions?

- When can we expect a complete 15 or 20-minute protocol session for full scoring analysis?

- Do we have access to user profile data (age, sex, height, weight, conditions) to enable normative scoring?

- Dataset 2 showed O2 at 8-9% with all clinical sensors at zero — is this a known firmware failure mode or a device wear issue?

# 8. Files Committed to kenisha_dev Branch

| File | Description |
| --- | --- |
| aerocardia_tcc_cleaned.csv | Cleaned Dataset 1 — TCC calibration run (360 rows) |
| aerocardia_user_session_cleaned.csv | Cleaned Dataset 3 — Real user session (26,312 rows) |
| kenisha_dev_findings.docx | This findings document |

AeroCardia DS Internship | Week 3 Updated Findings | Analyst: Kenisha | Branch: kenisha_dev